## Fluence Calculations
| Date of this Version | 2014-12-12 | [14 empty] | Determination of the "Petri Factor"
Germicidal Fluence (UV Dose) Calculations for a Low Pressure UV Lamp
[7 empty] | Programmed by Jim Bolton - Bolton Photosciences Inc., 628 Cheriton Cres., NW, Edmonton, AB, Canada T6R 2M5 |  |  |  | Steps:
[7 empty] | Tel: 780-439-4709; Fax: 780-439-7792; Email: jbolton@boltonuv.com |  |  |  | 1. Draw a 0.5 cm x 0.5 cm grid and place the center of the grid at the center of the collimated beam
[7 empty] | Comments and/or questions are welcome |  |  |  | 2. Measure the UV irradiance (with a radiometer) every 0.5 cm in the x and y directions and place the readings into columns S and X below. It is only
necessary to take readings out to 0.5 cm beyond the radius of the Petri dish (or other container to be used). The units do not matter since ratios are
calculated.
[7 empty] | DO NOT CHANGE ANY CELLS OTHER THAN THE CELLS WITH A YELLOW BACKGROUND |  |  |  | 3. The Petri factor is calculated in Cell CW151 and is also displayed in Cell K34.
INSTRUCTIONS AND NOTES
1. Set up a 'quasi' collimated beam apparatus. A 'collimating tube' may be used, but it is preferable to use circular 'masks' to define the beam. Make sure that | [10 empty] | UV Distribution across the Petri Dish for a low pressure UV Lamp
safety measures are taken to protect workers from exposure to the UV from the lamp. EYE PROTECTION IS AN ABSOLUTE REQUIREMENT.
2. Place the detector head of the UV radiometer on a horizontal surface, containing a 0.5 cm x 0.5 cm grid, such that the 'calibration plane' (see the Calibration | [12 empty] | Meter |  |  |  |  | Meter
Sheet provided by the manufacturer of the Radiometer) is at the level of where the top of the solution will be during exposures to the UV. | [10 empty] | x | y | Reading | Ratio |  | x | y | Reading | Ratio | [77 empty] | x | y | Ratio | x | y | Ratio
3. Determine the Petri Factor using the procedure given starting in Column Q.
4. Make sure that the spectrophotometer is balanced with distilled water in the same cuvette. Then measure the absorption coefficient (1 cm absorbance) | [10 empty] | 0 | -10 | 0 | 0 |  | -10 | 0 | 0 | 0 | [77 empty] | 0 | -10 | 0 | -10 | 0 | 0
at 254 nm for the water to be irradiated and insert into Cell G37. | [10 empty] | 0 | -9.5 | 0 | 0 |  | -9.5 | 0 | 0 | 0 | [77 empty] | 0 | -9.75 | 0 | -9.75 | 0 | 0
5. Insert the solution volume added to the Petri dish into Cell G31. | [10 empty] | 0 | -9 | 0 | 0 |  | -9 | 0 | 0 | 0 | [77 empty] | 0 | -9.5 | 0 | -9.5 | 0 | 0
6. Insert the stirrer volume into Cell G32. | [10 empty] | 0 | -8.5 | 0 | 0 |  | -8.5 | 0 | 0 | 0 | [77 empty] | 0 | -9.25 | 0 | -9.25 | 0 | 0
7. Insert the internal Petri dish diameter into Cell G33. | [10 empty] | 0 | -8 | 0 | 0 |  | -8 | 0 | 0 | 0 | [77 empty] | 0 | -9 | 0 | -9 | 0 | 0
8. Insert the distance from the center of the UV lamp to the surface of the water in the Petri Dish into cell G36. | [10 empty] | 0 | -7.5 | 0 | 0 |  | -7.5 | 0 | 0 | 0 | [77 empty] | 0 | -8.75 | 0 | -8.75 | 0 | 0
9. Insert the center meter reading into cell G40. | [10 empty] | 0 | -7 | 0 | 0 |  | -7 | 0 | 0 | 0 | [77 empty] | 0 | -8.5 | 0 | -8.5 | 0 | 0
10. Insert the desired Fluences (UV Doses) into cells E46 to E52. | [10 empty] | 0 | -6.5 | 0 | 0 |  | -6.5 | 0 | 0 | 0 | [77 empty] | 0 | -8.25 | 0 | -8.25 | 0 | 0
11. Remove the radiometer detector head and place a Petri Dish (or other container), containing the cell suspension, on a stirring motor placed so that the top of | [10 empty] | 0 | -6 | 0 | 0 |  | -6 | 0 | 0 | 0 | [77 empty] | 0 | -8 | 0 | -8 | 0 | 0
the solution is at the same level as that of the calibration plane of the detector head. Add a very small stir bar and make sure that the stirring rate is such that | [10 empty] | 0 | -5.5 | 0 | 0 |  | -5.5 | 0 | 0 | 0 | [77 empty] | 0 | -7.75 | 0 | -7.75 | 0 | 0
there is no vortex. | [10 empty] | 0 | -5 | 0 | 0 |  | -5 | 0 | 0 | 0 | [77 empty] | 0 | -7.5 | 0 | -7.5 | 0 | 0
12. Expose samples in the UV beam for the times calculated in rows 46 to 52. | [10 empty] | 0 | -4.5 | 0 | 0 |  | -4.5 | 0 | 0 | 0 | [77 empty] | 0 | -7.25 | 0 | -7.25 | 0 | 0
13. For a control, or zero dose determination, a sample solution is added to the Petri dish, set into the collimated beam apparatus and left stirring for a | [10 empty] | 0 | -4 | 91 | 0.91 |  | -4 | 0 | 91 | 0.91 | [77 empty] | 0 | -7 | 0 | -7 | 0 | 0
time equal to the average of the exposure times used. The resulting counts are used as the 0 dose concentration | [10 empty] | 0 | -3.5 | 93 | 0.93 |  | -3.5 | 0 | 93 | 0.93 | [77 empty] | 0 | -6.75 | 0 | -6.75 | 0 | 0
[11 empty] | 0 | -3 | 95 | 0.95 |  | -3 | 0 | 95 | 0.95 | [77 empty] | 0 | -6.5 | 0 | -6.5 | 0 | 0
[5 empty] | solution volume added to Petri dish = | 25 | mL |  | Water Factor = | 0.99173 | 0 | -2.5 | 97 | 0.97 |  | -2.5 | 0 | 97 | 0.97 | [77 empty] | 0 | -6.25 | 0 | -6.25 | 0 | 0
[5 empty] | stirrer volume = | 0.54 | mL |  | Divergence Factor = | 0.97077 | 0 | -2 | 97 | 0.97 |  | -2 | 0 | 97 | 0.97 | [77 empty] | 0 | -6 | 0 | -6 | 0 | 0
[5 empty] | Petri dish internal diameter = | 6 | cm |  | Reflection Factor = | 0.975 | 0 | -1.5 | 98 | 0.98 |  | -1.5 | 0 | 98 | 0.98 | [77 empty] | 0 | -5.75 | 0 | -5.75 | 0 | 0
[5 empty] | total volume in Petri dish = | 25.54 | mL |  | Petri Factor = | 0.96602 | 0 | -1 | 99 | 0.99 |  | -1 | 0 | 99 | 0.99 | [77 empty] | 0 | -5.5 | 0 | -5.5 | 0 | 0
[5 empty] | water path length = | 0.90329 | cm |  |  |  | 0 | -0.5 | 99 | 0.99 |  | -0.5 | 0 | 99 | 0.99 | [77 empty] | 0 | -5.25 | 0 | -5.25 | 0 | 0
[5 empty] | distance from UV lamp to top of water surface = | 30 | cm |  |  |  | 0 | 0 | 100 | 1 |  | 0 | 0 | 100 | 1 | [77 empty] | 0 | -5 | 0 | -5 | 0 | 0
[5 empty] | absorption coefficient at 254 nm = | 0.008 | cm-1 |  |  |  | 0 | 0.5 | 99 | 0.99 |  | 0.5 | 0 | 99 | 0.99 | [77 empty] | 0 | -4.75 | 0 | -4.75 | 0 | 0
[5 empty] | total absorbance at 254 nm = | 0.0072263 |  |  |  |  | 0 | 1 | 99 | 0.99 |  | 1 | 0 | 99 | 0.99 | [77 empty] | 0 | -4.5 | 0 | -4.5 | 0 | 0
[11 empty] | 0 | 1.5 | 98 | 0.98 |  | 1.5 | 0 | 98 | 0.98 | [77 empty] | 0 | -4.25 | 0.455 | -4.25 | 0 | 0.455
[5 empty] | Radiometer reading at the center of Petri Dish = | 0.22 | mW/cm2 |  |  |  | 0 | 2 | 97 | 0.97 |  | 2 | 0 | 97 | 0.97 | [77 empty] | 0 | -4 | 0.91 | -4 | 0 | 0.91
[5 empty] | True irradiance across the Petri dish = | 0.21252 | mW/cm2 |  |  |  | 0 | 2.5 | 97 | 0.97 |  | 2.5 | 0 | 97 | 0.97 | [77 empty] | 0 | -3.75 | 0.92 | -3.75 | 0 | 0.92
[11 empty] | 0 | 3 | 95 | 0.95 |  | 3 | 0 | 95 | 0.95 | [77 empty] | 0 | -3.5 | 0.93 | -3.5 | 0 | 0.93
[5 empty] | Average Germicidal Irradiance throughout the water volume = | 0.19949 | mW/cm2 |  |  |  | 0 | 3.5 | 93 | 0.93 |  | 3.5 | 0 | 93 | 0.93 | [77 empty] | 0 | -3.25 | 0.94 | -3.25 | 0 | 0.94
[11 empty] | 0 | 4 | 91 | 0.91 |  | 4 | 0 | 91 | 0.91 | [77 empty] | 0 | -3 | 0.95 | -3 | 0 | 0.95
|  |  | Time for a Fluence (UV Dose) of | 1 | mJ/cm2 = | 5.0128 | s | min | seconds |  | 0 | 4.5 | 0 | 0 |  | 4.5 | 0 | 0 | 0 | [77 empty] | 0 | -2.75 | 0.96 | -2.75 | 0 | 0.96
|  |  | Time for a Fluence (UV Dose) of | 10 | mJ/cm2 = | 50.1279 | s = | 0 | 50.1279 |  | 0 | 5 | 0 | 0 |  | 5 | 0 | 0 | 0 | [77 empty] | 0 | -2.5 | 0.97 | -2.5 | 0 | 0.97
|  |  | Time for a Fluence (UV Dose) of | 20 | mJ/cm2 = | 100.2559 | s = | 1 | 40.2559 |  | 0 | 5.5 | 0 | 0 |  | 5.5 | 0 | 0 | 0 | [77 empty] | 0 | -2.25 | 0.97 | -2.25 | 0 | 0.97
|  |  | Time for a Fluence (UV Dose) of | 30 | mJ/cm2 = | 150.3838 | s = | 2 | 30.3838 |  | 0 | 6 | 0 | 0 |  | 6 | 0 | 0 | 0 | [77 empty] | 0 | -2 | 0.97 | -2 | 0 | 0.97
|  |  | Time for a Fluence (UV Dose) of | 40 | mJ/cm2 = | 200.5117 | s = | 3 | 20.5117 |  | 0 | 6.5 | 0 | 0 |  | 6.5 | 0 | 0 | 0 | [77 empty] | 0 | -1.75 | 0.975 | -1.75 | 0 | 0.975
|  |  | Time for a Fluence (UV Dose) of | 50 | mJ/cm2 = | 250.6397 | s = | 4 | 10.6397 |  | 0 | 7 | 0 | 0 |  | 7 | 0 | 0 | 0 | [77 empty] | 0 | -1.5 | 0.98 | -1.5 | 0 | 0.98
|  |  | Time for a Fluence (UV Dose) of | 60 | mJ/cm2 = | 300.7676 | s = | 5 | 0.7676 |  | 0 | 7.5 | 0 | 0 |  | 7.5 | 0 | 0 | 0 | [77 empty] | 0 | -1.25 | 0.985 | -1.25 | 0 | 0.985
|  |  | Time for a Fluence (UV Dose) of | 70 | mJ/cm2 = | 350.8955 | s = | 5 | 50.8955 |  | 0 | 8 | 0 | 0 |  | 8 | 0 | 0 | 0 | [77 empty] | 0 | -1 | 0.99 | -1 | 0 | 0.99
[11 empty] | 0 | 8.5 | 0 | 0 |  | 8.5 | 0 | 0 | 0 | [77 empty] | 0 | -0.75 | 0.99 | -0.75 | 0 | 0.99
| Note: the exposure times should be at least 30 s. If they are calculated to be shorter, arrange the irradiation platform further away from the UV lamp so | [9 empty] | 0 | 9 | 0 | 0 |  | 9 | 0 | 0 | 0 | [77 empty] | 0 | -0.5 | 0.99 | -0.5 | 0 | 0.99
| that the irradiance will be smaller. If this is done, the Petri factor must be re-measured | [9 empty] | 0 | 9.5 | 0 | 0 |  | 9.5 | 0 | 0 | 0 | [77 empty] | 0 | -0.25 | 0.995 | -0.25 | 0 | 0.995
[11 empty] | 0 | 10 | 0 | 0 |  | 10 | 0 | 0 | 0 | [77 empty] | 0 | 0 | 1 | 0 | 0 | 1
[97 empty] | 0 | 0.25 | 0.995 | 0.25 | 0 | 0.995
[97 empty] | 0 | 0.5 | 0.99 | 0.5 | 0 | 0.99
[97 empty] | 0 | 0.75 | 0.99 | 0.75 | 0 | 0.99
[97 empty] | 0 | 1 | 0.99 | 1 | 0 | 0.99
[97 empty] | 0 | 1.25 | 0.985 | 1.25 | 0 | 0.985
[97 empty] | 0 | 1.5 | 0.98 | 1.5 | 0 | 0.98
[97 empty] | 0 | 1.75 | 0.975 | 1.75 | 0 | 0.975
[11 empty] | x |  | -10 | -9.75 | -9.5 | -9.25 | -9 | -8.75 | -8.5 | -8.25 | -8 | -7.75 | -7.5 | -7.25 | -7 | -6.75 | -6.5 | -6.25 | -6 | -5.75 | -5.5 | -5.25 | -5 | -4.75 | -4.5 | -4.25 | -4 | -3.75 | -3.5 | -3.25 | -3 | -2.75 | -2.5 | -2.25 | -2 | -1.75 | -1.5 | -1.25 | -1 | -0.75 | -0.5 | -0.25 | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.25 | 1.5 | 1.75 | 2 | 2.25 | 2.5 | 2.75 | 3 | 3.25 | 3.5 | 3.75 | 4 | 4.25 | 4.5 | 4.75 | 5 | 5.25 | 5.5 | 5.75 | 6 | 6.25 | 6.5 | 6.75 | 7 | 7.25 | 7.5 | 7.75 | 8 | 8.25 | 8.5 | 8.75 | 9 | 9.25 | 9.5 | 9.75 | 10 |  |  |  | 0 | 2 | 0.97 | 2 | 0 | 0.97
[11 empty] | y | [85 empty] | 0 | 2.25 | 0.97 | 2.25 | 0 | 0.97
[11 empty] | -10 | [83 empty] | 0 | 0 | 0 | 2.5 | 0.97 | 2.5 | 0 | 0.97
[11 empty] | -9.75 | [83 empty] | 0 | 0 | 0 | 2.75 | 0.96 | 2.75 | 0 | 0.96
[11 empty] | -9.5 | [83 empty] | 0 | 0 | 0 | 3 | 0.95 | 3 | 0 | 0.95
[11 empty] | -9.25 | [83 empty] | 0 | 0 | 0 | 3.25 | 0.94 | 3.25 | 0 | 0.94
[11 empty] | -9 | [83 empty] | 0 | 0 | 0 | 3.5 | 0.93 | 3.5 | 0 | 0.93
[11 empty] | -8.75 | [83 empty] | 0 | 0 | 0 | 3.75 | 0.92 | 3.75 | 0 | 0.92
[11 empty] | -8.5 | [83 empty] | 0 | 0 | 0 | 4 | 0.91 | 4 | 0 | 0.91
[11 empty] | -8.25 | [83 empty] | 0 | 0 | 0 | 4.25 | 0.455 | 4.25 | 0 | 0.455
[11 empty] | -8 | [83 empty] | 0 | 0 | 0 | 4.5 | 0 | 4.5 | 0 | 0
[11 empty] | -7.75 | [83 empty] | 0 | 0 | 0 | 4.75 | 0 | 4.75 | 0 | 0
[11 empty] | -7.5 | [83 empty] | 0 | 0 | 0 | 5 | 0 | 5 | 0 | 0
[11 empty] | -7.25 | [83 empty] | 0 | 0 | 0 | 5.25 | 0 | 5.25 | 0 | 0
[11 empty] | -7 | [83 empty] | 0 | 0 | 0 | 5.5 | 0 | 5.5 | 0 | 0
[11 empty] | -6.75 | [83 empty] | 0 | 0 | 0 | 5.75 | 0 | 5.75 | 0 | 0
[11 empty] | -6.5 | [83 empty] | 0 | 0 | 0 | 6 | 0 | 6 | 0 | 0
[11 empty] | -6.25 | [83 empty] | 0 | 0 | 0 | 6.25 | 0 | 6.25 | 0 | 0
[11 empty] | -6 | [83 empty] | 0 | 0 | 0 | 6.5 | 0 | 6.5 | 0 | 0
[11 empty] | -5.75 | [83 empty] | 0 | 0 | 0 | 6.75 | 0 | 6.75 | 0 | 0
[11 empty] | -5.5 | [83 empty] | 0 | 0 | 0 | 7 | 0 | 7 | 0 | 0
[11 empty] | -5.25 | [83 empty] | 0 | 0 | 0 | 7.25 | 0 | 7.25 | 0 | 0
[11 empty] | -5 | [83 empty] | 0 | 0 | 0 | 7.5 | 0 | 7.5 | 0 | 0
[11 empty] | -4.75 | [83 empty] | 0 | 0 | 0 | 7.75 | 0 | 7.75 | 0 | 0
[11 empty] | -4.5 | [83 empty] | 0 | 0 | 0 | 8 | 0 | 8 | 0 | 0
[11 empty] | -4.25 | [83 empty] | 0 | 0 | 0 | 8.25 | 0 | 8.25 | 0 | 0
[11 empty] | -4 | [83 empty] | 0 | 0 | 0 | 8.5 | 0 | 8.5 | 0 | 0
[11 empty] | -3.75 | [83 empty] | 0 | 0 | 0 | 8.75 | 0 | 8.75 | 0 | 0
[11 empty] | -3.5 | [83 empty] | 0 | 0 | 0 | 9 | 0 | 9 | 0 | 0
[11 empty] | -3.25 | [83 empty] | 0 | 0 | 0 | 9.25 | 0 | 9.25 | 0 | 0
[11 empty] | -3 | [41 empty] | 0.95 | [41 empty] | 0.95 | 1 | 0 | 9.5 | 0 | 9.5 | 0 | 0
[11 empty] | -2.75 | [37 empty] | 0.9504 | 0.9504 | 0.9504 | 0.9552 | 0.96 | 0.9552 | 0.9504 | 0.9504 | 0.9504 | [37 empty] | 8.5728 | 9 | 0 | 9.75 | 0 | 9.75 | 0 | 0
[11 empty] | -2.5 | [35 empty] | 0.9506 | 0.95545 | 0.9603 | 0.9603 | 0.9603 | 0.96515 | 0.97 | 0.96515 | 0.9603 | 0.9603 | 0.9603 | 0.95545 | 0.9506 | [35 empty] | 12.4742 | 13 | 0 | 10 | 0 | 10 | 0 | 0
[11 empty] | -2.25 | [34 empty] | 0.94575 | 0.9506 | 0.95545 | 0.9603 | 0.9603 | 0.9603 | 0.96515 | 0.97 | 0.96515 | 0.9603 | 0.9603 | 0.9603 | 0.95545 | 0.9506 | 0.94575 | [34 empty] | 14.3657 | 15
[11 empty] | -2 | [33 empty] | 0.9409 | 0.94575 | 0.9506 | 0.95545 | 0.9603 | 0.9603 | 0.9603 | 0.96515 | 0.97 | 0.96515 | 0.9603 | 0.9603 | 0.9603 | 0.95545 | 0.9506 | 0.94575 | 0.9409 | [33 empty] | 16.2475 | 17
[11 empty] | -1.75 | [32 empty] | 0.94575 | 0.94575 | 0.95062 | 0.9555 | 0.96037 | 0.96525 | 0.96525 | 0.96525 | 0.97013 | 0.975 | 0.97013 | 0.96525 | 0.96525 | 0.96525 | 0.96037 | 0.9555 | 0.95062 | 0.94575 | 0.94575 | [32 empty] | 18.2227 | 19
[11 empty] | -1.5 | [31 empty] | 0.9506 | 0.9506 | 0.9506 | 0.9555 | 0.9604 | 0.9653 | 0.9702 | 0.9702 | 0.9702 | 0.9751 | 0.98 | 0.9751 | 0.9702 | 0.9702 | 0.9702 | 0.9653 | 0.9604 | 0.9555 | 0.9506 | 0.9506 | 0.9506 | [31 empty] | 20.2174 | 21
[11 empty] | -1.25 | [31 empty] | 0.95545 | 0.95545 | 0.95545 | 0.96037 | 0.9653 | 0.97023 | 0.97515 | 0.97515 | 0.97515 | 0.98008 | 0.985 | 0.98008 | 0.97515 | 0.97515 | 0.97515 | 0.97023 | 0.9653 | 0.96037 | 0.95545 | 0.95545 | 0.95545 | [31 empty] | 20.3205 | 21
[11 empty] | -1 | [30 empty] | 0.9504 | 0.9603 | 0.9603 | 0.9603 | 0.96525 | 0.9702 | 0.97515 | 0.9801 | 0.9801 | 0.9801 | 0.98505 | 0.99 | 0.98505 | 0.9801 | 0.9801 | 0.9801 | 0.97515 | 0.9702 | 0.96525 | 0.9603 | 0.9603 | 0.9603 | 0.9504 | [30 empty] | 22.3245 | 23
[11 empty] | -0.75 | [30 empty] | 0.9504 | 0.9603 | 0.9603 | 0.9603 | 0.96525 | 0.9702 | 0.97515 | 0.9801 | 0.9801 | 0.9801 | 0.98505 | 0.99 | 0.98505 | 0.9801 | 0.9801 | 0.9801 | 0.97515 | 0.9702 | 0.96525 | 0.9603 | 0.9603 | 0.9603 | 0.9504 | [30 empty] | 22.3245 | 23
[11 empty] | -0.5 | [30 empty] | 0.9504 | 0.9603 | 0.9603 | 0.9603 | 0.96525 | 0.9702 | 0.97515 | 0.9801 | 0.9801 | 0.9801 | 0.98505 | 0.99 | 0.98505 | 0.9801 | 0.9801 | 0.9801 | 0.97515 | 0.9702 | 0.96525 | 0.9603 | 0.9603 | 0.9603 | 0.9504 | [30 empty] | 22.3245 | 23
[11 empty] | -0.25 | [30 empty] | 0.9552 | 0.96515 | 0.96515 | 0.96515 | 0.97013 | 0.9751 | 0.98008 | 0.98505 | 0.98505 | 0.98505 | 0.99003 | 0.995 | 0.99003 | 0.98505 | 0.98505 | 0.98505 | 0.98008 | 0.9751 | 0.97013 | 0.96515 | 0.96515 | 0.96515 | 0.9552 | [30 empty] | 22.4373 | 23
[11 empty] | 0 | [29 empty] | 0.95 | 0.96 | 0.97 | 0.97 | 0.97 | 0.975 | 0.98 | 0.985 | 0.99 | 0.99 | 0.99 | 0.995 | 1 | 0.995 | 0.99 | 0.99 | 0.99 | 0.985 | 0.98 | 0.975 | 0.97 | 0.97 | 0.97 | 0.96 | 0.95 | [29 empty] | 24.45 | 25
[11 empty] | 0.25 | [30 empty] | 0.9552 | 0.96515 | 0.96515 | 0.96515 | 0.97013 | 0.9751 | 0.98008 | 0.98505 | 0.98505 | 0.98505 | 0.99003 | 0.995 | 0.99003 | 0.98505 | 0.98505 | 0.98505 | 0.98008 | 0.9751 | 0.97013 | 0.96515 | 0.96515 | 0.96515 | 0.9552 | [30 empty] | 22.4373 | 23
[11 empty] | 0.5 | [30 empty] | 0.9504 | 0.9603 | 0.9603 | 0.9603 | 0.96525 | 0.9702 | 0.97515 | 0.9801 | 0.9801 | 0.9801 | 0.98505 | 0.99 | 0.98505 | 0.9801 | 0.9801 | 0.9801 | 0.97515 | 0.9702 | 0.96525 | 0.9603 | 0.9603 | 0.9603 | 0.9504 | [30 empty] | 22.3245 | 23
[11 empty] | 0.75 | [30 empty] | 0.9504 | 0.9603 | 0.9603 | 0.9603 | 0.96525 | 0.9702 | 0.97515 | 0.9801 | 0.9801 | 0.9801 | 0.98505 | 0.99 | 0.98505 | 0.9801 | 0.9801 | 0.9801 | 0.97515 | 0.9702 | 0.96525 | 0.9603 | 0.9603 | 0.9603 | 0.9504 | [30 empty] | 22.3245 | 23
[11 empty] | 1 | [30 empty] | 0.9504 | 0.9603 | 0.9603 | 0.9603 | 0.96525 | 0.9702 | 0.97515 | 0.9801 | 0.9801 | 0.9801 | 0.98505 | 0.99 | 0.98505 | 0.9801 | 0.9801 | 0.9801 | 0.97515 | 0.9702 | 0.96525 | 0.9603 | 0.9603 | 0.9603 | 0.9504 | [30 empty] | 22.3245 | 23
[11 empty] | 1.25 | [31 empty] | 0.95545 | 0.95545 | 0.95545 | 0.96037 | 0.9653 | 0.97023 | 0.97515 | 0.97515 | 0.97515 | 0.98008 | 0.985 | 0.98008 | 0.97515 | 0.97515 | 0.97515 | 0.97023 | 0.9653 | 0.96037 | 0.95545 | 0.95545 | 0.95545 | [31 empty] | 20.3205 | 21
[11 empty] | 1.5 | [31 empty] | 0.9506 | 0.9506 | 0.9506 | 0.9555 | 0.9604 | 0.9653 | 0.9702 | 0.9702 | 0.9702 | 0.9751 | 0.98 | 0.9751 | 0.9702 | 0.9702 | 0.9702 | 0.9653 | 0.9604 | 0.9555 | 0.9506 | 0.9506 | 0.9506 | [31 empty] | 20.2174 | 21
[11 empty] | 1.75 | [32 empty] | 0.94575 | 0.94575 | 0.95062 | 0.9555 | 0.96037 | 0.96525 | 0.96525 | 0.96525 | 0.97013 | 0.975 | 0.97013 | 0.96525 | 0.96525 | 0.96525 | 0.96037 | 0.9555 | 0.95062 | 0.94575 | 0.94575 | [32 empty] | 18.2227 | 19
[11 empty] | 2 | [33 empty] | 0.9409 | 0.94575 | 0.9506 | 0.95545 | 0.9603 | 0.9603 | 0.9603 | 0.96515 | 0.97 | 0.96515 | 0.9603 | 0.9603 | 0.9603 | 0.95545 | 0.9506 | 0.94575 | 0.9409 | [33 empty] | 16.2475 | 17
[11 empty] | 2.25 | [34 empty] | 0.94575 | 0.9506 | 0.95545 | 0.9603 | 0.9603 | 0.9603 | 0.96515 | 0.97 | 0.96515 | 0.9603 | 0.9603 | 0.9603 | 0.95545 | 0.9506 | 0.94575 | [34 empty] | 14.3657 | 15
[11 empty] | 2.5 | [35 empty] | 0.9506 | 0.95545 | 0.9603 | 0.9603 | 0.9603 | 0.96515 | 0.97 | 0.96515 | 0.9603 | 0.9603 | 0.9603 | 0.95545 | 0.9506 | [35 empty] | 12.4742 | 13
[11 empty] | 2.75 | [37 empty] | 0.9504 | 0.9504 | 0.9504 | 0.9552 | 0.96 | 0.9552 | 0.9504 | 0.9504 | 0.9504 | [37 empty] | 8.5728 | 9
[11 empty] | 3 | [41 empty] | 0.95 | [41 empty] | 0.95 | 1
[11 empty] | 3.25 | [83 empty] | 0 | 0
[11 empty] | 3.5 | [83 empty] | 0 | 0
[11 empty] | 3.75 | [83 empty] | 0 | 0
[11 empty] | 4 | [83 empty] | 0 | 0
[11 empty] | 4.25 | [83 empty] | 0 | 0
[11 empty] | 4.5 | [83 empty] | 0 | 0
[11 empty] | 4.75 | [83 empty] | 0 | 0
[11 empty] | 5 | [83 empty] | 0 | 0
[11 empty] | 5.25 | [83 empty] | 0 | 0
[11 empty] | 5.5 | [83 empty] | 0 | 0
[11 empty] | 5.75 | [83 empty] | 0 | 0
[11 empty] | 6 | [83 empty] | 0 | 0
[11 empty] | 6.25 | [83 empty] | 0 | 0
[11 empty] | 6.5 | [83 empty] | 0 | 0
[11 empty] | 6.75 | [83 empty] | 0 | 0
[11 empty] | 7 | [83 empty] | 0 | 0
[11 empty] | 7.25 | [83 empty] | 0 | 0
[11 empty] | 7.5 | [83 empty] | 0 | 0
[11 empty] | 7.75 | [83 empty] | 0 | 0
[11 empty] | 8 | [83 empty] | 0 | 0
[11 empty] | 8.25 | [83 empty] | 0 | 0
[11 empty] | 8.5 | [83 empty] | 0 | 0
[11 empty] | 8.75 | [83 empty] | 0 | 0
[11 empty] | 9 | [83 empty] | 0 | 0
[11 empty] | 9.25 | [83 empty] | 0 | 0
[11 empty] | 9.5 | [83 empty] | 0 | 0
[11 empty] | 9.75 | [83 empty] | 0 | 0
[11 empty] | 10 | [83 empty] | 0 | 0
[95 empty] | 426.0133 | 441
[94 empty] | Petri factor = | 0.96602
Note on G33: Jim Bolton: Measure carefully with a plastic mm ruler or better with a pair of calipers.
Note on G35: Jim Bolton: The water path length may be entered directly.